## Slide 1: British Thoracic Society Thoracic Ultrasound Physics Of Ultrasound
Dr Uzoma Nnajiuba
Radiologist

## Slide 2: Objectives
Describe the behavior of sound waves through different densities i.e. air, fluid, solid tissue, consolidated lung.
Explain how different probe frequencies influence sonographic appearances and relationship to structure/pathology.
Recognise and explain normal artefacts, specifically A-lines, B-lines and mirror imaging.
Explain the Doppler effect and its relevance to TUS.
Describe the rationale and practical reasons to vary gain/ depth/focal point.

## Slide 3: Ultrasound waves
Sound waves that are at a higher frequency than that which is audible (>20kHz).
Typically 3-15 MHz in diagnostic imaging.
  High frequency allows them to be pulsed in beam
Non-ionizing

## Slide 4: Piezoelectric Effect
When a mechanical strain is applied across certain crystals, they generate a voltage. Piezoelectric effect (Detection).
When a voltage is applied across these crystals they become mechanically strained (generate a mechanical force) Reverse Piezoelectric effect (Production).
Lead Zirconate Titanate (PZT)

## Slide 5: Echo Production
In US echoes produced when waves meet an interface between two materials with different acoustic impedances (z).
  Similar to how you see structures on a CXR because they have different radiographic densities.

## Slide 6: Echo Production
A.I. is a measure of how difficult it is for US waves to travel through a medium.
  Density.
  Speed of sound within material.
Greater the difference in the ‘z’ the stronger the echo.
First interface is transducer crystals and air.
  Need coupling gel
AIR
X
Soft Tissue
Soft Tissue
Gel

## Slide 7: Image Formation in Thoracic US
ST: 1.58
Muscle: 1.7
Bone: 6
Lung: 0.18
Air: 0.0004
A Line

## Slide 8: A lines
Move (with pleural line) during ventilation.
Present?
Yes
Air deep to pleura
No
No movement.
Aerated Lung
Air + Solid or liquid
-Pulmonary Oedema
-Consolidation
-Fibrosis
Pneumothorax

## Slide 9: Image Formation: B Lines
>2 = Pathological

## Slide 10: B Lines
A lines

## Slide 11: Dense consolidation
If there is even more liquid/solid content in the lung then will actually get echoes emitted from deeper within the lung.

## Slide 12: Dense consolidation

## Slide 13: Pleural Effusion

## Slide 14: Mirror Image artefact
Reverberation again
Notes: So I’ll try to illustrate that with this diagram.
As you can see on the right of the screen you have this liver tumor which is reflected in the right hemidiaphragm and looks like its also in the right lung base.
That occurs because you have a nicely aerated lung and so a highly reflective soft tissue air interface at the right hemidiaphragm (click)
So what happens is US waves are emitted and travel from the transducer to the hemidiaphragm and there they are essentially all reflected back to the transducer (and you’ll see a bright line as the hemidiaphragm)
Because the relfection is so strong they travel back as a powerful beam and behave almost as if they’ve just been emitted and so on the way back they bounce off structures (click) within the liver and go back to the hemidiaphragm before then returning back to the transducer.
The TD doesn’t know that this second echo is a reverberation that has been from the TD to the diaphragm back off a structure in the liver then back to the diaphragm and then back to the TD. It just assumes it’s a new echo that has taken twice as long to return so it places that structure further away i.e. deep to the diaphragm and in that way strcutures in the liver can be reflected in the diaphragm as evidenced by this diagram where you see a liver lesion reflected in the diaphragm.

## Slide 15: Pleural Effusion
Lung
Pleural Fluid
Diaphragm
Vertebra
Liver/Spleen
Posterior Acoustic Enhancement

## Slide 16: Transducer (Probe) Types
Linear
Curvilinear

## Slide 17: Linear Probes
Higher Frequency 7.5 – 10 MHz (short wavelength)
  Higher resolution
  Poorer penetration
Rectangular FOV width of the probe
Small superficial structures
  Thyroids
  Testes
  Breast
  MSK
Subtle pathology
Can’t scan deep structures

## Slide 18: Curvilinear Probes
Lower Frequency 2-4 MHz (Long wavelength)
  Lower resolution
  Higher penetration
Sector shaped FOV
Larger deep structures
  HPB
  KUB
  Thoracic

## Slide 19: Doppler Effect
When emitted US waves interact with a moving interface, their frequency changes.
Can be displayed as colour.
Could use to distinguish between debris (non-vascular) and pleural thickening (vascular)
Color Doppler: Directional
Pulse Doppler: Non-Directional

## Slide 20: Optimizing Image (Knobology)
Depth
Gain
Focus

## Slide 21: Gain
Adjusting the gain changes the brightness of the image.
Usually controlled by a knob
Gain should be adjusted so that fluid appears black and soft tissue is mid-gray.
Just Right
Too bright
Too Dark

## Slide 22: Depth
The depth of the field of view.
Displayed in cm on the side of the screen.
Structure being scanned should fill middle of the screen.
Correct
Too Deep

## Slide 23: Focal Point
Point where US beam is narrowest, and resolution is highest.
Usually adjusted by a dial and indicated by a marker on the side of the screen.
Should be focused on the object of interest.
Correct
Incorrect Focus

## Slide 24: Summary
Describe the behaviour of sound waves through different densities i.e. air, fluid, solid tissue, consolidated lung.
Explain how different probe frequencies influence sonographic appearances and relationship to structure/pathology.
Recognise and explain normal artefacts, specifically A-lines, B-lines and mirror imaging.
Explain the Doppler effect and its relevance to TUS.
Describe the rationale and practical reasons to vary gain/depth/focal point.

## Slide 25: Thank you

## Slide 26: MCQs
US coupling gel is used to:
  Provide lubrication when scanning
  Make the probe head warmer to improve patient comfort while scanning
  Cool the transducer crystals
  Prevent complete reflection of US waves as they leave the transducer crystals
  Increase the resolution of the image

## Slide 27: MCQs
A lines are a reverberation artefact produced at the interface between:
  Ribs and the pleura
  Transducer crystals and the skin.
  Fluid and air within the alveoli
  Pleura and aerated lung
  Transducer crystals and the air

## Slide 28: MCQs
The mirror image artefact in the context of:
  Normally aerated lung
  Pleural effusion
  Dense consolidation
  Pulmonary oedema
  Pulmonary Fibrosis

## Slide 29: MCQs
Adjusting the gain changes the:
  Size of the image
  Brightness of the image
  Resolution of the image
  Focal point
  Frame rate

## Slide 30: MCQs
Curvilinear probes are:
  Higher frequency and used for deeper structures
  Higher frequency and used for superficial structures
  Colour doppler probes used to detect blood flow
  Lower frequency and used for superficial structures
  Lower frequency and used for deeper structures